## Slide 1
Pooled sputum testing
for Tuberculosis diagnosis

## Slide 2
After this training, participants will be able to:
Understand when pooled testing is appropriate
Describe the pooled testing workflow
Prepare and test pooled samples
Manage positive pools and deconvolution testing
Ensure quality assurance and traceability
Objectives
of the Training

## Slide 3
Expands access to molecular TB testing
Uses fewer cartridges
Reduces laboratory workload
Supports large screening activities
Why Pooled
Testing Matters

## Slide 4
Interpretation of Results
If pooled result = MTB not detected → all specimens are interpreted as negative → no further testing is required
If pooled result = MTB detected or trace → perform individual testing for all specimens in the pool → use remaining processed specimen If pooled result = Error, invalid, or no result (non-valid) Follow the Laboratory procedure for non-valid results. → Repeat pooled test if volume remains, OR → Perform individual testing of all specimens in the pool. Clinical decisions are based on the individual test results.
What is Pooled Testing

## Slide 5
When to Use
Pooled Testing
Key conditions:
Low to moderate positivity (<30%)
Adequate sputum volume
Ability to perform reflex testing
Typical settings:
Community screening
High throughput laboratories
Outreach campaigns

## Slide 6
When NOT to
Use Pooled Testing
❌ High TB positivity settings
❌ Insufficient sputum volume
❌ Invalid samples such as extrapulmonary samples
❌ Need for immediate individual result (e.g. risk of drug-resistant TB)
❌ Risk of paucibacillary TB including children, people living with HIV, and those with prior TB treatment

## Slide 7
Sputum Sample Flow for Pooled Testing
MTB Detected OR Not Detected
Individual Xpert Ultra assay
Invalid/Error/No results
Repeat
2 mL of mixture
Report result to clinical team
INDIVIDUAL TESTING
Standard of Care
Sputum (≥2 mL)
Add 2 parts of SR to 1 part of sputum
Total volume sputum /
SR mixture: ≥6 mL
Optional: Take a swab for smear microscopy
Pooled Xpert Ultra assay
MTB Detected OR Not Detected
Invalid/Error/No results
Repeat
4 consecutive samples available
Create pools of 4 samples
(1 mL from each sample in pool)
Fridge available
< 4 samples available
Store samples overnight
Pool remaining samples
(1.5 mL from each sample in pool)
Additional samples available to complete pools of 4 on following day
Yes
No
Record result
POOLED TESTING
Shake 10 to 20 times or vortex for at least 10 seconds for full homogenisation
2 mL of pooled mixture
Fridge unavailable

## Slide 8
Specimen Eligibility
Receive specimen
2. Check specimen eligibility
Sputum (not saliva)
Adequate quality
Minimum volume (> 1 mL)
Correctly labelled

## Slide 9
Preparing Samples
Arrange specimens in a muffin tray so samples for one pool are in the same row.
Add an empty container in line to serve as the pooled sample container.
For pools of four, place four containers in adjacent wells.
This helps organise specimens and prevents mixing errors.

## Slide 10
Preparing Samples
Process specimen with sample reagent according to manufacturer’s instructions:
A.
10-20 times
B.
C.
E.
10-20 times
D.

## Slide 11
Forming the Pool
Label pool container
Mix the specimen-reagent mixture briefly to re-homogenise
Transfer an equal aliquot using a new sterile pipette or pipette tip
Combine aliquots into the pool container
Mix pooled sample
Example of preparing a ~4 mL pool of four samples

## Slide 12
After aliquoting all specimens, gently mix the pooled sample to ensure uniformity.
Using a sterile pipette, transfer the required volume of the pooled sample into the cartridge or reaction tube and run the test on the LC-aNAAT platform following standard procedures.
Where possible, prepare pools and load cartridges in separate areas to reduce the risk of cross contamination.
Running the Pooled Test

## Slide 13
Interpretation
of Pooled Results
Pooled Result | Action Required | Final Interpretation
MTB not detected | No further testing required | All individuals considered negative according to the molecular test result
MTB detected | Perform individual testing of all specimens in the pool | Individual results determine final diagnosis
MTB detected (trace) | Perform individual testing of all specimens in the pool | Individual results determine final diagnosis
Error / invalid / no result | Repeat pooled test or proceed to individual testing according to laboratory procedures | Based on repeat or individual results

## Slide 14
If the pooled test result is MTB detected (including trace), each individual specimen included in the pool must undergo deconvolution individual testing:
Retrieve all processed individual specimens that contributed to the pool.
Mix each specimen again briefly to re-homogenise.
Perform individual LC-aNAAT testing following the standard testing procedure.
Reflex Individual Testing

## Slide 15
Preventing Errors
and Contamination
Key points:
New pipette for each sample
Equal volumes
Keep containers closed
Clean workspace
Separate samples (individual and pooled) and cartridge loading areas

## Slide 16
Traceability and
Sample Tracking
Each specimen must have a unique individual identifier.
Each pool must be assigned a unique Pool ID.
Pool IDs must be linked to all specimens included in that pool to allow traceability and individual retesting if needed.
Pool ID | Specimen ID 1 | Specimen ID 2 | Specimen ID 3 | Specimen ID 4
LIV-20260309-P01 | TB-001245 | TB-001246 | TB-001247 | TB-001248
Example:

## Slide 17
Recording of
Pooled Test Results
Minimum information to record in the laboratory register or electronic laboratory information system (LIS) used for routine LC-aNAAT testing:
Pool ID
Date of testing
Pool size
Individual specimen IDs included in the pool
Individual test result when applicable
Pooled test result (MTB trace detected, MTB not detected, error, invalid, or no result), initial test result and repeated, if any
Linking pooled and individual samples:
Where possible, add a Pool ID column in the laboratory register.
If registers cannot be modified, use a separate pooled testing log sheet to track pools

## Slide 18
Final results must always be reported at the individual specimen level for clinical management as in routine individual testing:
If the pooled test is MTB not detected, report each specimen as MTB not detected
If the pooled test is MTB detected (including trace), report individual results after reflex testing
If the pooled test is Error / invalid / no result, report individual results after repeat or reflex testing
Reporting of Results

## Slide 19
Monitoring unexpected results:
  - Positive pool but all individual tests negative
  - Unexpected increase in positive pools
  - Increase in non-valid results (Error / invalid / no result)
  - Failure to perform deconvolution testing of all specimens in positive pools
Monitor discordant pooled testing results:
  - Expected rate ~3 to 10% (positive pool, all individual MTB not detected), mainly with very low or trace results
  Rates outside this range or high / medium / low pooled results with all individual MTB not detected should trigger review and repeat testing
  Check: Specimen identification, pipetting, workspace organisation, and SOP adherence.
Quality Assurance
and Quality Control

## Slide 20
Document all non-valid results (Error / invalid / no result), procedural errors, and corrective actions. Corrective actions may include:
Repeating tests (if specimen volume remains)
Reviewing pooled testing procedures
Retraining staff on aliquoting and specimen identification
Checking instrument maintenance and calibration
Documentation
and Corrective Actions
Report major deviations to laboratory supervisors.

## Slide 21
Continue participation in LC-aNAAT EQA programmes.
Include pooled testing in:
Proficiency testing schemes
Supervisory visits and laboratory audits
NTP laboratory quality monitoring systems
External Quality
Assurance
Ensures pooled testing maintains diagnostic performance comparable to individual testing

## Slide 22
The Start4All Partnership

## Slide 23
The Start4All consortium expresses its thanks to all stakeholders, collaborators, and participants who supported our research efforts.
For more information, see our online Acknowledgements page: www.lstmed.ac.uk/start4all/acknowledgements
Start4All is funded by Unitaid
Acknowledgements

## Slide 24
lstmed.ac.uk/start4all